## Slide 1: SCD Calibration
Steven Tomczyk, Alfred deWijn
and the SCD Team
<14pt Times New Roman>
Date of Presentation

## Slide 2: Objective
Polarimetric calibration of the SCD Instrument
This formulation is based on:
seminal paper by del Toro Iniesta and Collados (dTI&C), “Optimum modulation and demodulation matrices for solar polarimetry”, App. Opt., 39, 1637, 2000
and presentation by deWijn on “The Concept of Optimal Calibration”
The notation of dTI&C will be used here. The matrix dimensions are row x column

## Slide 3: Theory
The measurement process is defined by Imeas = O Iin
Iin is the Stokes vector (I,Q,U,V) input to the polarimeter (4 vector)
O is the modulation matrix (n x 4 matrix)
n is the number of modulation states
Imeas is the vector of measured intensities (n vector)
For SCD, n = 4

## Slide 4: Theory
The goal is to obtain Iin from Imeas
This process is called demodulation
We can define Iin = D Imeas
Where D is the demodulation matrix (4 x n)
dTI&C find that it is not feasible in general to obtain D by taking the inverse of O

## Slide 5: Theory
Instead dTI&C compute D from O by first computing the matrix A = OT O (4 x 4 matrix)
Then D = A-1 OT = (OT O)-1 OT
The matrix A also contains valuable information about the efficiency of the polarimeter which relates to the noise on the demodulated Stokes vector (see dTI&C)
The efficiencies are computed from the diagonal elements of A by ei = 1/(nA-1i,i)

## Slide 6: Calibration
Computation of D requires knowledge of O which is obtained by the calibration process
For SCD calibration, we will assume that we know the properties (Mueller matrix vs. ) of the calibration optics
The  dependence is important since the SCD was designed to have high polarimetric efficiency over the entire wavelength range, but D varies greatly with  (see Tomczyk etal, JOSA, 2010)
We will also (initially) assume that the light entering the polarimeter is unpolarized with Stokes vector [1,0,0,0]

## Slide 7: Calibration
Following deWijn we can draw an analogy to the dTI&C procedure and define a calibration matrix
C = [Iin,1 Iin,2 … Iin,m], (4 x m matrix) where m is the number of calibration states and Iin,j is the Stokes vector produced by calibration state j (for SCD, m = 6)
Iin,j = MMj IC , where IC is the Stokes vector of light entering the calibration optics and MMj is the Mueller matrix of calibration state j (we are assuming IC = [1,0,0,0] initially)
The m sets of measured intensities for the calibration are then Imeas,j = O Iin,j or Imeas n,m = O C

## Slide 8: Calibration
where Imeas n,m is a n row by m column matrix containing the n intensities for the m calibration states
deWijn defines a matrix E (m x 4) such that O = Imeas n,m E
then O = O C E and C E = 1.
E is the nominal inverse of C computed by E = CT (C CT)-1 (the diagonal terms of C CT provide calibration efficiencies)
and O can be computed by O = Imeas n,m E

## Slide 9: Calibration Procedure
create C = [Iin,1 Iin,2 … Iin,m] matrix (4 x m) from the calibration optics Mueller matrices using Iin,j = MMj IC , assuming IC = [1,0,0,0]
compute the E matrix (m x 4) from the C matrix using E = CT (C CT)-1
create the Imeas n,m matrix (n x m) from the calibration observations comprised of the n intensities for the m calibration states (after bias and flat correction)
compute the O matrix (n x 4) from O = Imeas n,m E
compute A = OT O, then D = A-1 OT

## Slide 10: Consistency Check
It is possible to check the consistency of the calibrations using the additional clear observed as part of the calibration scheme
Apply the D matrix to the clear observation
It should yield a Stokes vector very close to [1,0,0,0]
If it does not, substitute the clear Stokes vector for Ic , go back to step 1 and recompute D
Iterate if needed

## Slide 11: Apply
Now that D has been computed, it can be applied to the measurements
Bias subtract and flat field divide the intensity vector Imeas at each pixel
Apply the D matrix to obtain the Stokes vector
Note that D may vary across the image and may need to be computed at various spatial points and interpolated
D will need to be computed for each wavelength region but should not vary appreciably across a spectral line

## Slide 12: Next Steps
I recommend that 1 main program and 5 subroutines be written to perform SCD calibration
The 5 subroutines will perform the 5 corresponding tasks on the Calibration Procedure slide
Steve will write:
    SCD_Create_Cmatrix
    SCD_Compute_Ematrix
    SCD_Compute_Omatrix
    SCD_Compute_Dmatrix
    That correspond to steps 1, 2, 4 and 5 respectively
AISAS will write SCD_Get_Cal_Data that will perform step 3

## Slide 13: Next Steps
AISAS will write the SCD_Calibration main program that will call all of these routines
My goal is to have a draft of the 4 routines by the end of next week
All of this is up for discussion
Comments and feedback are very welcome
It is strongly recommended that all involved become very familiar with the dTI&C paper